## Procedure
TESS Video Board v6.1 Bring-Up Procedure Results
| Unit Under Test (Must be Video Board version 6.1 or 6.1T): |  |  |  | Notes in red must be addressed before this procedure can be considered ready.
| P/N: Video 6.1 PCB Rev 1 | Serial Number: 6.1-T Vid P02 |  |  | Need expected values for most steps.
Test Engineer: Ed Bokhour
Step | Test Description | Expected Value | Measured Value | Accepted | Notes | Initials | Date
BEFORE POWER UP:
1 | Verify work area is ESD safe. |  |  | OK |  | EB | 2015-09-17
2 | Photograph assembly, both sides. Store photo in sheet 5 of this spreadsheet (Assembly Photos (5). |  |  | OK |  | EB | 2015-09-17
3 | Weigh assembly. Note any non-flight configurations here. |  | 0.125 kg | OK | No conformal coating. | EB | 2015-09-17
4 | Note test equipment P/N's and S/N's (e.g. DMMs, oscilloscopes, thermal imaging unit, power supplies). |  |  | OK | DMM: Fluke 117 sn 29940033WS. Calibration status unknown. Scope: Tek TDS 3054C, s/n TSS2005044 (property of E. Bokhour). Calibration out of date. Observatory Simulator: label = 37-518. | EB | 2015-09-17
5 | Perform visual inspection under stereo microscope. Check for workmanship and mechanical damage. | No mechanical damage. Workmanship to IPC standard (TBD). |  | OK |  | EB | 2015-09-17
6 | Verify no missing components. |  | None missing | OK | AirBorn connectors not installed yet. | EB | 2015-09-17
7 | Verify any DNP/DNI parts are not installed. | X1/X1/R7,X1/X2/R7,X1/X3/R7,X1/X4/R7,X2/X1/R7,X2/X2/R7,X2/X3/R7,X2/X4/R7,X3/X1/R7,X3/X2/R7,X3/X3/R7,X3/X4/R7,X4/X1/R7,X4/X2/R7,X4/X3/R7,X4/X4/R7,X1/X1/C18,X1/X2/C18,X1/X3/C18,X1/X4/C18,X2/X1/C18,X2/X2/C18,X2/X3/C18,X2/X4/C18,X3/X1/C18,X3/X2/C18,X3/X3/C18,X3/X4/C18,X4/X1/C18,X4/X2/C18,X4/X3/C18,X4/X4/C18 | All DNP parts are not installed. | OK |  | EB | 2015-09-17
8 | (Deleted)
9 | Verify correct orientation of all components (chips, diodes, polarized capacitors, etc.) | All polarized components and chips are in correct orientation. | All polarized parts match orientation on silkscreen. | OK |  | EB | 2015-09-17
10 | Verify chips installed are the correct parts. If specified in the design, verify date codes for relevant components. | All parts are correct type. | All parts look correct. Smaller devices verified by similarity to others with same part description in layout. | OK |  | EB | 2015-09-17
11 | Inspect connectors for damage, misalignment, bent pins. | None. | No damage. | OK | AirBorn connectors not installed. | EB | 2015-09-17
12 | Examine through holes for all four AirBorn connectors, J1, J2, J3, J4, for clearance. Use backlighting with stereo microscope to make sure through holes are clear of residual solder or other debris, and are sufficiently open to allow AirBorn connector pins to pass through without mechanical resistance. | Holes are clear for entire depth and full diameter of barrel. | Already Installed | OK |  | EB | 2015-09-17
13 | Install required test headers and note locations. | None required | None installed. | OK |  | EB | 2015-09-17
| With power OFF and no external connections to PCB, measure resistance on all power lines. All measurements referenced to circuit GND unless otherwise noted. |  | These values inform the reference set.
14 | GND, multiple locations | <=0.1 Ohms | 0.1 Ohms, typ | OK | Multiple locations on J5 (stack connector); one side of all C3's, C28's, and C30's. | EB | 2015-09-17
15 | +5V | >2M Ohm | >2.3M | OK | J5:180 | EB | 2015-09-17
16 | +15V | 59K Ohm | 59.14K | OK | J5:176 | EB | 2015-09-17
17 | -12V | 43.9K Ohm | 43.77K | OK | J5:178 | EB | 2015-09-17
18 | -50V | >2M Ohm | 2.94M Ohms | OK | PP/C13.2 (or PP/R12.2) | EB | 2015-09-17
19 | +24V | 57.5K Ohms | 57.49K Ohms | OK | PP/C14.2 | EB | 2015-09-17
20 | VS12 | 21K Ohms | 20.98K Ohms | OK | Q2.2 to U15.4 | EB | 2015-09-17
21 | VS34 | 21K Ohms | 21.00K Ohms | OK | Q3.2 to U15.4 | EB | 2015-09-17
22 | +3.3Dac (X1) | 18K Ohms | 18.37K | OK | X1/U7.8 or U7.9 | EB | 2015-09-17
22.1 | +3.3Dac (X2) | 18K Ohms | 18.34K | OK | X2/U7.8 or U7.9 | EB | 2015-09-17
23 | +3.3Dac (X3) | 18K Ohms | 18.51K | OK | X3/U9.8 or U9.9 | EB | 2015-09-17
23.1 | +3.3Dac (X4) | 18K Ohms | 18.41K | OK | X4/U7.8 or U7.9 | EB | 2015-09-17
With power OFF, measure resistance of all connector pins to circuit GND.
24 | GND pins | 0 +/- 0.1 Ohms |  |  | See Data Tables (2), Table 2
25 | Signal pins | R values match reference set |  |  | See Data Tables (2), Table 2
26 | No-connect pins | Open circuit |  |  | See Data Tables (2), Table 2
27 | (Deleted)
PREP FOR POWER ON TESTS
28 | Verify work area is ESD safe
29 | Put test setup in documented configuration. This will include a 5.3 Interface Board and power supply stack. Also LSE computer running CameraLink software (Aoshiro) with the appropriate software to run the stack. |  |  | OK
30 | Mate Driver Board to Interface Board and verify the proper connection. Record the version and serial number of the Interface Board in the measured value cell.
31 | If testing flight hardware, verify any required connector savers are installed. | None required for rev 6.1 (non-flight)
32 | Attach power and CameraLink cables to the Interface Board.
POWER ON TESTS
33 | Turn ON Observatory Simulator (rear panel power switch first, then right-hand front panel switch to right-most position).
34 | Inspect Video PCB with thermal imaging unit to check for abnormally hot spots. | Normal heating profile |  |  | Compare to Reference Thermal Images (sheet 3).
35 | Record all supply currents and voltages in Data Table 1 (sheet 2). Verify voltages and currents are within 1% of expected vales. | See Table 1, sheet 2. |  |  | % tolerance needs to be verified. Use power supply front panel for current measurements.
36 | (Deleted)
Measure voltage on all open connector pins with power ON
37 | (Deleted)
38 | GND pins | 0 +/- 0.01V
39 | Signal pins | Voltages match reference set. |  |  | Document full set of voltage measurements for one (verified good) PCB. Compare other boards to this set.
40 | No-connect pins | 0 +/- 0.00 V
41 | Power OFF.
| Test programming of FPGA. |  |  |  | NOTE: Files which must be up to date: FPE_Wrapper.bin, and Artix_FPE_MemConfig_tmp.xls. Compile memory data from spreadsheet by running ./exceltobin_tmp.py (in memfiles/scripts). This generates 7 *.asc and 7 *.bin files used by ./LoadFPE, and must be run before ./LoadFPE.sh.
42 | Observatory Simulator ON. Verify nominal supply currents. | 290 mA on +5V, 10 mA on +15V, 56mA on -12V
43 | On GSE computer, open a terminal window and type "rlwrap nc -u 192.168.100.1 5555". This will be referred to as the "NC window". Hit "return" twice to get prompt. | Prompt "FPE1>" should be displayed.
44 | In NC terminal window, type "cam_status". | Terminal should return 0x00e0000d or 0x00e0000e.
45 | Open a second terminal window, and navigate to /observatorysimulator/examples
46 | In second terminal window, type "./LoadFPE.sh" | Both windows should display status messages confirming successful loading of FPGA software and program memories.
47 | In second terminal window, navigate to /opt/Data/TESS/FPE6.1. |  |  |  | use /tmp instead?
48 | Create a new directory with today's date as the name. Name format should be YYYYMMDD, as in 20150810 for Aug 10th, 2015.
49 | Create a readme file in this new directory, describing today's tests.
50 | In the NC window, type "cam_start_frames". | Frame data should be displayed.
51 | In the second terminal window, type "tess_obssim" to start catching frames. Hit <ctrl C> to exit after at least one frame.
52 | Examine frame using ds9 software. Store image in Reference Raw Image (4) of this spreadsheet. | Similar to Reference Raw Image, sheet 4.
53 | In second terminal window, type "readHK >hkdata.txt" to capture only housekeeping data and capture it in a file named hkdata.txt. | File includes only HK data.
54 | Compare data in hkdata.txt to reference set in sheet 6, Housekeeping Data. | Data matches reference set within TBD% tolerance.
55 | Power OFF

## Data Tables (2)
STEP 35, TABLE 1: POWER SUPPLY VOLTAGES AND CURRENTS
| +5V | +15V | -12V | +3.3Dac
Measurement location: | J5:180 | J5:176 | J5:178 | X1_X10_R1.1
Voltage
Expected Values | 5 | 15 | -12 | 3.3
Error | -100% | -100% | -100% | -100%
Current |  |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Expected Values | 0.29 | 0.01 | 0.056 | N/A | N/A | N/A | N/A | N/A | N/A
Error | -100% | -100% | -100%
STEP 25, TABLE 2a: |  |  |  |  | STEPS 38-40, TABLE 3
Resistance (power off, no external connections, wait time 30 sec or less for capacitive charging) |  |  |  |  | Voltage Measurements, power ON, power and data cables to Obs Sim.
All resistances referenced to J6 pin A1 (meter COM = pin A1) |  |  |  |  | All voltages referenced to power connector pin 1 (COM = pin 1)
Connector | Pin | Signal Name | Expected Value (Ohms, +/-10%) | Measured Value (Ohms) | Expected Value (Volts, +/- 5%) | Measured Value (Volts)
J6 | A1 | GND | 0.2 | 0.2 | 0
J6 | A2 | SDO-A-3 | >3 M | 3.44M | 0.049
J6 | A3 | GND | 0.2 | 0.2 | 0
J6 | A4 | SDO-B-3 | >3 M | 3.44M | 0.051
J6 | A5 | GND | 0.2 | 0.2 | 0
J6 | A6 | SDO-C-3 | >3 M | 3.46M | 0.048
J6 | A7 | GND | 0.2 | 0.2 | 0
J6 | A8 | SDO-D-3 | >3 M | 3.46M | 0.05
J6 | A9 | GND | 0.2 | 0.2 | 0
J6 | A10 | CNV-3 | >3 M | 3.43M | 0.05
J6 | A11 | GND | 0.2 | 0.2 | 0
J6 | A12 | SCK-3 | >3 M | 3.5M | 0.05
J6 | A13 | GND | 0.2 | 0.2 | 0
J6 | A14 | GND | 0.2 | 0.2 | 0.05
J6 | A15 | GND | 0.2 | 0.2 | 0
J6 | A16 | HKA5 | >3 M | 9.79M | 0.05
J6 | A17 | HKA3 | >3 M | 9.74M | 0
J6 | A18 | DD | >3 M | 20.65M | 0.05
J6 | A19 | GND | 0.2 | 0.2 | 0
J6 | A20 | GND | 0.2 | 0.2 | 0.07
J6 | A21 | GND | 0.2 | 0.2 | 0
J6 | A22 | SDO-A-4 | >3 M | 3.42M | 0.05
J6 | A23 | GND | 0.2 | 0.2 | 0
J6 | A24 | SDO-B-4 | >3 M | 0.2 | 0.05
J6 | A25 | GND | 0.2 | 0.2 | 0
J6 | A26 | SDO-C-4 | >3 M | 3.43M | 0.07
J6 | A27 | GND | 0.2 | 0.2 | 0
J6 | A28 | SDO-D-4 | >3 M | 3.44M | 0.07
J6 | A29 | GND | 0.2 | 0.2 | 0
J6 | A30 | CNV-4 | >3 M | 3.47M | 0.07
J6 | A31 | GND | 0.2 | 0.2 | 0
J6 | A32 | SCK-4 | >3 M | 3.47M | 0.07
J6 | A33 | GND | 0.2 | 0.2 | 0
J6 | A34 | -12 | >1 M | 3.04M | 0.07
J6 | B1 | P1-FS-3 | 0L | 0L | 0
J6 | B2 | P2-FS-3 | 0L | 0L | 0.07
J6 | B3 | P3-FS-3 | 0L | 0L | 0
J6 | B4 | RG-3 | 0L | 0L | 0.07
J6 | B5 | P1-IA-3 | 0L | 0L | 0
J6 | B6 | P2-IA-3 | 0L | 0L | 0.07
J6 | B7 | P3-IA-3 | 0L | 0L | 0
J6 | B8 | P3-OR-3 | 0L | 0L | 0.05
J6 | B9 | P2-OR-3 | 0L | 0L | 0
J6 | B10 | P1-OR-3 | 0L | 0L | 0.05
J6 | B11 | ID-3 | 0L | 0L | 0
J6 | B12 | GND | 0.2 | 0.2 | 0.05
J6 | B13 | GND | 0.2 | 0.2 | 0
J6 | B14 | DCK | >3 M | 18.48M | 0.05
J6 | B15 | HKA6 | >3 M | 9.75M | 0
J6 | B16 | HKA4 | >3 M | 9.87M | 0.05
J6 | B17 | HKA2 | >3 M | 10M | 0
J6 | B18 | HKA1 | >3 M | 16.55M | 0
J6 | B19 | HKA0 | >3 M | 17.53M | 0
J6 | B20 | GND | 0.2 | 0.2 | 4.706
J6 | B21 | 2.5 | 24-28K | 24.16K | 4.706
J6 | B22 | P1-FS-4 | 0L | 0L | 4.706
J6 | B23 | P2-FS-4 | 0L | 0L | 14.7
J6 | B24 | P3-FS-4 | 0L | 0L | 0
J6 | B25 | RG-4 | 0L | 0L | 0
J6 | B26 | P1-IA-4 | 0L | 0L | 0
J6 | B27 | P2-IA-4 | 0L | 0L | 4.706
J6 | B28 | P3-IA-4 | 0L | 0L | 4.706
J6 | B29 | P3-OR-4 | 0L | 0L | 4.706
J6 | B30 | P2-OR-4 | 0L | 0L | 0
J6 | B31 | P1-OR-4 | 0L | 0L | 0.06
J6 | B32 | ID-4 | 0L | 0L | 0.06
J6 | B33 | 15 | >3 M | 7.81M | 0.06
J6 | B34 | 5 | >3 M | 3.25M | 0.06
J8 | A1 | GND | 0.2 | 0.2 | 0.06
J8 | A2 | SCK-1 | >3 M | 3.46M | 0.06
J8 | A3 | GND | 0.2 | 0.2 | 0.07
J8 | A4 | CNV-1 | >3 M | 3.46M | 0.07
J8 | A5 | GND | 0.2 | 0.2 | 1.31
J8 | A6 | SDO-A-1 | >3 M | 3.45M | 0.06
J8 | A7 | GND | 0.2 | 0.2 | 0.07
J8 | A8 | SDO-B-1 | >3 M | 3.47M | 0.07
J8 | A9 | GND | 0.2 | 0.2 | 0.07
J8 | A10 | SDO-C-1 | >3 M | 3.47M | 0.07
J8 | A11 | GND | 0.2 | 0.2 | 0.07
J8 | A12 | SDO-D-1 | >3 M | 3.47 | 0.07
J8 | A13 | GND | 0.2 | 0.2 | 0.07
J8 | A14 | GND | 0.2 | 0.2 | 0.06
J8 | A15 | GND | 0.2 | 0.2 | 0.06
J8 | A16 | GND | 0.2 | 0.2 | 0.06
J8 | A17 | Int | >3 M | 9.17M | 0.06
J8 | A18 | GND | 0.2 | 0.2 | 0.06
J8 | A19 | \_Hold\_ | >3 M | 39.45M | 0
J8 | A20 | GND | 0.2 | 0.2 | 0
J8 | A21 | Clamp | >3 M | 4.957M | 0
J8 | A22 | GND | 0.2 | 0.2 | 4.707
J8 | A23 | GND | 0.2 | 0.2 | 4.707
J8 | A24 | SCK-2 | 388K | 3.43M | 4.707
J8 | A25 | GND | 0.2 | 0.2 | 14.69
J8 | A26 | CNV-2 | >3 M | 3.44M | 0
J8 | A27 | GND | 0.2 | 0.2 | 0
J8 | A28 | SDO-A-2 | >3 M | 3.44M | 0
J8 | A29 | GND | 0.2 | 0.2 | 4.707
J8 | A30 | SDO-B-2 | >3 M | 3.43M | 4.707
J8 | A31 | GND | 0.2 | 0.2 | 4.707
J8 | A32 | SDO-C-2 | >3 M | 3.43M | 0
J8 | A33 | GND | 0.2 | 0.2 | 0
J8 | A34 | SDO-D-2 | >3 M | 3.45M | 4.705
J8 | B1 | 24 | >3 M | 4.41M | 4.705
J8 | B2 | -50 | >3 M | 3.098M | 4.705
J8 | B3 | P1-FS-1 | 0L | 0L | 0
J8 | B4 | P2-FS-1 | 0L | 0L | 0
J8 | B5 | P3-FS-1 | 0L | 0L | 0
J8 | B6 | P3-OR-1 | 0L | 0L | 14.69
J8 | B7 | P2-OR-1 | 0L | 0L | 4.705
J8 | B8 | P1-OR-1 | 0L | 0L | 4.705
J8 | B9 | RG-1 | 0L | 0L | 4.705
J8 | B10 | P1-IA-1 | 0L | 0L | 0
J8 | B11 | P2-IA-1 | 0L | 0L | 0
J8 | B12 | P3-IA-1 | 0L | 0L | 0
J8 | B13 | ID-1 | 0L | 0L | 0.6
J8 | B14 | GND | 0.2 | 0.2 | 0.6
J8 | B15 | DCS1 | >3 M | 9.12M | 0.6
J8 | B16 | DCS0 | >3 M | 4.92M | 0.6
J8 | B17 | DCS2 | >3 M | 9.22 | 0.6
J8 | B18 | DCS3 | >3 M | 8.75M | 0.6
J8 | B19 | HKCOM | >3 M | 8.2M | 0.6
J8 | B20 | TCOM | 0L | 0L | 0.6
J8 | B21 | RTDCOM | >3 M | 13.22M | 0.6
J8 | B22 | GND | 0.2 | 0.2 | 0.7
J8 | B23 | GND | 0.2 | 0.2 | 0.7
J8 | B24 | P1-FS-2 | 0L | 0L | 0.6
J8 | B25 | P2-FS-2 | 0L | 0L | 0.7
J8 | B26 | P3-FS-2 | 0L | 0L | 0.7
J8 | B27 | P3-OR-2 | 0L | 0L | 0.7
J8 | B28 | P2-OR-2 | 0L | 0L | 0
J8 | B29 | P1-OR-2 | 0L | 0L | 0.6
J8 | B30 | RG-2 | 0L | 0L | 0.6
J8 | B31 | P1-IA-2 | 0L | 0L | 0.6
J8 | B32 | P2-IA-2 | 0L | 0L | 0.6
J8 | B33 | P3-IA-2 | 0L | 0L | 0.6
J8 | B34 | ID-2 | 0L | 0L | 0.6
STEP 25, TABLE 2b:

## Reference Thermal Images (3)
REFERENCE IMAGES | CAPTURED IMAGES, THIS UUT

## Reference Raw Image (4)
REFERENCE IMAGE | IMAGE TAKEN WITH THIS UUT

## Housekeeping Data (6)
Housekeeping Data
REFERENCE VALUES (from IB sn03, 9/10/15, "readHK" command) | Captured Data for this UUT | Reference signal names, cam_hsk command
image 0 125956343.240028 125956345.607440 125956343.240031 2.367412 2.367409
HK index 0, raw 0x0000, 0 |  | [(0, 'ccd1_output_gate'),
HK index 1, raw 0x0017, 23 |  | (1, 'ccd1_input_gate_1'),
HK index 2, raw 0x0004, 4 |  | (2, 'ccd1_input_gate_2'),
HK index 3, raw 0x0006, 6 |  | (3, 'ccd1_scupper'),
HK index 4, raw 0x0007, 7 |  | (4, 'ccd1_reset_drain'),
HK index 5, raw 0x0005, 5 |  | (5, 'ccd1_backside'),
HK index 6, raw 0x0004, 4 |  | (6, 'ccd1_substrate'),
HK index 7, raw 0x0005, 5 |  | (7, 'ccd1_board_temperature'),
HK index 8, raw 0x0006, 6 |  | (8, 'ccd1_output_drain_a'),
HK index 9, raw 0x0005, 5 |  | (9, 'ccd1_output_drain_b'),
HK index 10, raw 0x0006, 6 |  | (10, 'ccd1_output_drain_c'),
HK index 11, raw 0x0007, 7 |  | (11, 'ccd1_output_drain_d'),
HK index 12, raw 0x0007, 7 |  | (12, 'ccd1_output_source_a'),
HK index 13, raw 0x0005, 5 |  | (13, 'ccd1_output_source_b'),
HK index 14, raw 0x0005, 5 |  | (14, 'ccd1_output_source_c'),
HK index 15, raw 0x0005, 5 |  | (15, 'ccd1_output_source_d'),
HK index 16, raw 0x0004, 4 |  | (16, 'ccd2_output_gate'),
HK index 17, raw 0x0007, 7 |  | (17, 'ccd2_input_gate_1'),
HK index 18, raw 0x0004, 4 |  | (18, 'ccd2_input_gate_2'),
HK index 19, raw 0x0005, 5 |  | (19, 'ccd2_scupper'),
HK index 20, raw 0x0006, 6 |  | (20, 'ccd2_reset_drain'),
HK index 21, raw 0x0004, 4 |  | (21, 'ccd2_backside'),
HK index 22, raw 0x0006, 6 |  | (22, 'ccd2_substrate'),
HK index 23, raw 0x0005, 5 |  | (23, 'ccd2_board_temperature'),
HK index 24, raw 0x0005, 5 |  | (24, 'ccd2_output_drain_a'),
HK index 25, raw 0x0005, 5 |  | (25, 'ccd2_output_drain_b'),
HK index 26, raw 0x0005, 5 |  | (26, 'ccd2_output_drain_c'),
HK index 27, raw 0x0005, 5 |  | (27, 'ccd2_output_drain_d'),
HK index 28, raw 0x0006, 6 |  | (28, 'ccd2_output_source_a'),
HK index 29, raw 0x0005, 5 |  | (29, 'ccd2_output_source_b'),
HK index 30, raw 0x0004, 4 |  | (30, 'ccd2_output_source_c'),
HK index 31, raw 0x0005, 5 |  | (31, 'ccd2_output_source_d'),
HK index 32, raw 0x0005, 5 |  | (32, 'ccd3_output_gate'),
HK index 33, raw 0x0006, 6 |  | (33, 'ccd3_input_gate_1'),
HK index 34, raw 0x1388, 5000 |  | (34, 'ccd3_input_gate_2'),
HK index 35, raw 0x0010, 16 |  | (35, 'ccd3_scupper'),
HK index 36, raw 0x0020, 32 |  | (36, 'ccd3_reset_drain'),
HK index 37, raw 0x0080, 128 |  | (37, 'ccd3_backside'),
HK index 38, raw 0x0000, 0 |  | (38, 'ccd3_substrate'),
HK index 39, raw 0x0001, 1 |  | (39, 'ccd3_board_temperature'),
HK index 40, raw 0x0000, 0 |  | (40, 'ccd3_output_drain_a'),
HK index 41, raw 0x0000, 0 |  | (41, 'ccd3_output_drain_b'),
HK index 42, raw 0x0000, 0 |  | (42, 'ccd3_output_drain_c'),
HK index 43, raw 0x0041, 65 |  | (43, 'ccd3_output_drain_d'),
HK index 44, raw 0x0000, 0 |  | (44, 'ccd3_output_source_a'),
HK index 45, raw 0x0000, 0 |  | (45, 'ccd3_output_source_b'),
HK index 46, raw 0x0000, 0 |  | (46, 'ccd3_output_source_c'),
HK index 47, raw 0x0000, 0 |  | (47, 'ccd3_output_source_d'),
HK index 48, raw 0x0000, 0 |  | (48, 'ccd4_output_gate'),
HK index 49, raw 0x0000, 0 |  | (49, 'ccd4_input_gate_1'),
HK index 50, raw 0x0000, 0 |  | (50, 'ccd4_input_gate_2'),
HK index 51, raw 0x0000, 0 |  | (51, 'ccd4_scupper'),
HK index 52, raw 0x0000, 0 |  | (52, 'ccd4_reset_drain'),
HK index 53, raw 0x0000, 0 |  | (53, 'ccd4_backside'),
HK index 54, raw 0x0000, 0 |  | (54, 'ccd4_substrate'),
HK index 55, raw 0x0000, 0 |  | (55, 'ccd4_board_temperature'),
HK index 56, raw 0x0000, 0 |  | (56, 'ccd4_output_drain_a'),
HK index 57, raw 0x0000, 0 |  | (57, 'ccd4_output_drain_b'),
HK index 58, raw 0x0000, 0 |  | (58, 'ccd4_output_drain_c'),
HK index 59, raw 0x0000, 0 |  | (59, 'ccd4_output_drain_d'),
HK index 60, raw 0x0000, 0 |  | (60, 'ccd4_output_source_a'),
HK index 61, raw 0x0000, 0 |  | (61, 'ccd4_output_source_b'),
HK index 62, raw 0x0000, 0 |  | (62, 'ccd4_output_source_c'),
HK index 63, raw 0x0000, 0 |  | (63, 'ccd4_output_source_d'),
HK index 64, raw 0x0020, 32 |  | (64, 'ccd1_parallel_high'),
HK index 65, raw 0x0017, 23 |  | (65, 'ccd1_parallel_low'),
HK index 66, raw 0x0005, 5 |  | (66, 'ccd1_serial_high'),
HK index 67, raw 0x0005, 5 |  | (67, 'ccd1_serial_low'),
HK index 68, raw 0x0005, 5 |  | (68, 'ccd1_reset_high'),
HK index 69, raw 0x0007, 7 |  | (69, 'ccd1_reset_low'),
HK index 70, raw 0x0005, 5 |  | (70, 'ccd1_input_diode_high'),
HK index 71, raw 0x0006, 6 |  | (71, 'ccd1_input_diode_low'),
HK index 72, raw 0x0006, 6 |  | (72, 'ccd2_parallel_high'),
HK index 73, raw 0x0007, 7 |  | (73, 'ccd2_parallel_low'),
HK index 74, raw 0x0005, 5 |  | (74, 'ccd2_serial_high'),
HK index 75, raw 0x0005, 5 |  | (75, 'ccd2_serial_low'),
HK index 76, raw 0x0005, 5 |  | (76, 'ccd2_reset_high'),
HK index 77, raw 0x0006, 6 |  | (77, 'ccd2_reset_low'),
HK index 78, raw 0x0006, 6 |  | (78, 'ccd2_input_diode_high'),
HK index 79, raw 0x0004, 4 |  | (79, 'ccd2_input_diode_low'),
HK index 80, raw 0x0006, 6 |  | (80, 'ccd3_parallel_high'),
HK index 81, raw 0x0005, 5 |  | (81, 'ccd3_parallel_low'),
HK index 82, raw 0x0005, 5 |  | (82, 'ccd3_serial_high'),
HK index 83, raw 0x0005, 5 |  | (83, 'ccd3_serial_low'),
HK index 84, raw 0x0006, 6 |  | (84, 'ccd3_reset_high'),
HK index 85, raw 0x0006, 6 |  | (85, 'ccd3_reset_low'),
HK index 86, raw 0x0005, 5 |  | (86, 'ccd3_input_diode_high'),
HK index 87, raw 0x0005, 5 |  | (87, 'ccd3_input_diode_low'),
HK index 88, raw 0x0006, 6 |  | (88, 'ccd4_parallel_high'),
HK index 89, raw 0x0006, 6 |  | (89, 'ccd4_parallel_low'),
HK index 90, raw 0x0007, 7 |  | (90, 'ccd4_serial_high'),
HK index 91, raw 0x0006, 6 |  | (91, 'ccd4_serial_low'),
HK index 92, raw 0x0005, 5 |  | (92, 'ccd4_reset_high'),
HK index 93, raw 0x0005, 5 |  | (93, 'ccd4_reset_low'),
HK index 94, raw 0x0004, 4 |  | (94, 'ccd4_input_diode_high'),
HK index 95, raw 0x0006, 6 |  | (95, 'ccd4_input_diode_low'),
HK index 96, raw 0x0006, 6 |  | (96, 'heater_1_current'),
HK index 97, raw 0x0006, 6 |  | (97, 'heater_2_current'),
HK index 98, raw 0x1388, 5000 |  | (98, 'heater_3_current'),
HK index 99, raw 0x0010, 16 |  | (99, None),
HK index 100, raw 0x0020, 32 |  | (100, None),
HK index 101, raw 0x0080, 128 |  | (101, None),
HK index 102, raw 0x0000, 0 |  | (102, None),
HK index 103, raw 0x0001, 1 |  | (103, None),
HK index 104, raw 0x0000, 0 |  | (104, None),
HK index 105, raw 0x0000, 0 |  | (105, '+15'),
HK index 106, raw 0x0000, 0 |  | (106, '+5'),
HK index 107, raw 0x0042, 66 |  | (107, '-12'),
HK index 108, raw 0x0000, 0 |  | (108, '+3.3f'),
HK index 109, raw 0x0000, 0 |  | (109, '+2.5f'),
HK index 110, raw 0x0000, 0 |  | (110, '+1.8f'),
HK index 111, raw 0x0000, 0 |  | (111, '+1f'),
HK index 112, raw 0x0000, 0 |  | (112, 'pt1000_sensor_1'),
HK index 113, raw 0x0000, 0 |  | (113, 'pt1000_sensor_2'),
HK index 114, raw 0x0000, 0 |  | (114, 'pt1000_sensor_3'),
HK index 115, raw 0x0000, 0 |  | (115, 'pt1000_sensor_4'),
HK index 116, raw 0x0000, 0 |  | (116, 'pt1000_sensor_5'),
HK index 117, raw 0x0000, 0 |  | (117, 'pt1000_sensor_6'),
HK index 118, raw 0x0000, 0 |  | (118, 'pt1000_sensor_7'),
HK index 119, raw 0x0000, 0 |  | (119, 'pt1000_sensor_8'),
HK index 120, raw 0x0000, 0 |  | (120, 'pt1000_sensor_9'),
HK index 121, raw 0x0000, 0 |  | (121, 'pt1000_sensor_10'),
HK index 122, raw 0x0000, 0 |  | (122, 'pt1000_sensor_11'),
HK index 123, raw 0x0000, 0 |  | (123, 'pt1000_sensor_12'),
HK index 124, raw 0x0000, 0 |  | (124, 'alcu_sensor_ccd1'),
HK index 125, raw 0x0000, 0 |  | (125, 'alcu_sensor_ccd2'),
HK index 126, raw 0x0000, 0 |  | (126, 'alcu_sensor_ccd3'),
HK index 127, raw 0x0000, 0 |  | (127, 'alcu_sensor_ccd4')]
HK index 128, raw 0x0040, 64 |  | (126, 'alcu_sensor_ccd3'),
HK index 129, raw 0x0017, 23 |  | (127, 'alcu_sensor_ccd4')]
HK index 130, raw 0x3dda, 15834
HK index 131, raw 0xf119, -3815
HK index 132, raw 0x1fbf, 8127
HK index 133, raw 0xe167, -7833
HK index 134, raw 0x17af, 6063
HK index 135, raw 0xd164, -11932
HK index 136, raw 0xa2ff, -23809
HK index 137, raw 0xa32d, -23763
HK index 138, raw 0x3dbb, 15803
HK index 139, raw 0xf0ff, -3841
HK index 140, raw 0x1f89, 8073
HK index 141, raw 0xe15e, -7842
HK index 142, raw 0x1996, 6550
HK index 143, raw 0xd331, -11471
HK index 144, raw 0xa388, -23672
HK index 145, raw 0xa33a, -23750
HK index 146, raw 0x0005, 5
HK index 147, raw 0x0004, 4
HK index 148, raw 0x0006, 6
HK index 149, raw 0x0003, 3
HK index 150, raw 0x0005, 5
HK index 151, raw 0x0005, 5
HK index 152, raw 0x0004, 4
HK index 153, raw 0x0004, 4
HK index 154, raw 0x0004, 4
HK index 155, raw 0x0005, 5
HK index 156, raw 0x0005, 5
HK index 157, raw 0x0006, 6
HK index 158, raw 0x0006, 6
HK index 159, raw 0x0006, 6
HK index 160, raw 0x0004, 4
HK index 161, raw 0x0007, 7
HK index 162, raw 0x1388, 5000
HK index 163, raw 0x0010, 16
HK index 164, raw 0x0020, 32
HK index 165, raw 0x0080, 128
HK index 166, raw 0x0000, 0
HK index 167, raw 0x0001, 1
HK index 168, raw 0x0000, 0
HK index 169, raw 0x0000, 0
HK index 170, raw 0x0000, 0
HK index 171, raw 0x0043, 67
HK index 172, raw 0x0000, 0
HK index 173, raw 0x0000, 0
HK index 174, raw 0x0000, 0
HK index 175, raw 0x0000, 0
HK index 176, raw 0x0000, 0
HK index 177, raw 0x0000, 0
HK index 178, raw 0x0000, 0
HK index 179, raw 0x0000, 0
HK index 180, raw 0x0000, 0
HK index 181, raw 0x0000, 0
HK index 182, raw 0x0000, 0
HK index 183, raw 0x0000, 0
HK index 184, raw 0x0000, 0
HK index 185, raw 0x0000, 0
HK index 186, raw 0x0000, 0
HK index 187, raw 0x0000, 0
HK index 188, raw 0x0000, 0
HK index 189, raw 0x0000, 0
HK index 190, raw 0x0000, 0
HK index 191, raw 0x0000, 0
HK index 192, raw 0x0060, 96
HK index 193, raw 0x0017, 23
HK index 194, raw 0x0004, 4
HK index 195, raw 0x0004, 4
HK index 196, raw 0xd99a, -9830
HK index 197, raw 0x8bac, -29780
HK index 198, raw 0xe637, -6601
HK index 199, raw 0x5da8, 23976
HK index 200, raw 0xf208, -3576
HK index 201, raw 0xec9a, -4966
HK index 202, raw 0x7fff, 32767
HK index 203, raw 0xf851, -1967
HK index 204, raw 0x7fff, 32767
HK index 205, raw 0x7fff, 32767
HK index 206, raw 0x7fff, 32767
HK index 207, raw 0x7fff, 32767
HK index 208, raw 0x7fff, 32767
HK index 209, raw 0x7fff, 32767
HK index 210, raw 0x7fff, 32767
HK index 211, raw 0x7fff, 32767
HK index 212, raw 0x7fff, 32767
HK index 213, raw 0x7fff, 32767
HK index 214, raw 0x7fff, 32767
HK index 215, raw 0x7fff, 32767
HK index 216, raw 0x7fff, 32767
HK index 217, raw 0x7fff, 32767
HK index 218, raw 0x0009, 9
HK index 219, raw 0x7fff, 32767
HK index 220, raw 0x0005, 5
HK index 221, raw 0x0006, 6
HK index 222, raw 0x0006, 6
HK index 223, raw 0x0004, 4
HK index 224, raw 0x0004, 4
HK index 225, raw 0x0005, 5
HK index 226, raw 0x1388, 5000
HK index 227, raw 0x0010, 16
HK index 228, raw 0x0020, 32
HK index 229, raw 0x0080, 128
HK index 230, raw 0x0000, 0
HK index 231, raw 0x0001, 1
HK index 232, raw 0x0000, 0
HK index 233, raw 0x0000, 0
HK index 234, raw 0x0000, 0
HK index 235, raw 0x0044, 68
HK index 236, raw 0x0000, 0
HK index 237, raw 0x0000, 0
HK index 238, raw 0x0000, 0
HK index 239, raw 0x0000, 0
HK index 240, raw 0x0000, 0
HK index 241, raw 0x0000, 0
HK index 242, raw 0x0000, 0
HK index 243, raw 0x0000, 0
HK index 244, raw 0x0000, 0
HK index 245, raw 0x0000, 0
HK index 246, raw 0x0000, 0
HK index 247, raw 0x0000, 0
HK index 248, raw 0x0000, 0
HK index 249, raw 0x0000, 0
HK index 250, raw 0x0000, 0
HK index 251, raw 0x0000, 0
HK index 252, raw 0x0000, 0
HK index 253, raw 0x0000, 0
HK index 254, raw 0x0000, 0
HK index 255, raw 0x0000, 0